Session Type

Physics

Topic

Beam Delivery and Nozzle Design

Presentation Preference

Oral

Abstract Number

PTC57-0077

Abstract Title

Innovative strip silicon detectors for proton beam monitoring: preliminary results.

Authors

A. Vignati1, O. Hammad Ali2, A. Attili1, M. Donetti3, F. Fausti1, S. Giordanengo1, L. Manganaro2, V. Monaco2, R. Sacchi2, R. Cirio2.
1National Institute for Nuclear Physics INFN, Turin division, Turin, Italy.
2National Institute for Nuclear Physics INFN- Università degli Studi di Torino, Physics Department, Turin, Italy.
3Fondazione CNAO, Medical Physics, Pavia, Italy.

Abstract Text

Unlike the legacy gas ionization chambers, solid state detectors offer large granularity and sensitivity to single protons and would ideally be suited for beam monitoring in therapy applications. However, signal pileup, radiation damage and the readout complexity prevented their use so far on high flux therapeutic beams. 
Innovative silicon low-gain avalanche detectors optimized for time resolution (Ultra Fast Silicon Detectors - UFSDs) were recently proposed, where sensors as thin as 50 µm provide signals of ~ 1ns time duration with time resolutions of tenths of ps, and large enough signal-to-noise ratio to efficiently discriminate the proton signal. 
The MoVeIT project of the INFN is investigating the use of this technology in Particle Therapy by developing two monitoring devices, one to directly count individual protons at high rates, the second to measure the beam energy with time-of-flight techniques. This requires the design of custom UFSD sensors as well VLSI readout electronics. 
From the simulations’ results and the first beam tests with UFSD pads, strip detectors were produced by FBK in Trento, with two geometries (30 mm and 15 mm length) and different doping modalities to improve radiation hardness. In parallel, prototypes of a new TERA10 readout chip have been submitted to the foundry.
The aim of this contribution is to review the advancement of the project and to report on the results of the tests of UFSD strip sensors with the therapeutic proton beam of CNAO.